## Slide 1
Stoyan Stoynev et al.
US Magnet Development Program
Fermi National Accelerator Laboratory
Strain gauges in the 15 T FNAL prototype
October 11, 2017

## Slide 2: Introduction
During the latest coil readiness review for layer 1 and 2 for the 15 T magnet at FNAL a discussion on strain gauges was advised
This presentation summarizes the information on strain gauges and intention of use
Limitations and considerations are emphasized
The slides are aimed to serve as a guide for the discussion

## Slide 3: Strain gauges at FNAL
Gauges suitable for
Stainless Steel
Skin, Bullets
(long)
Layer 3/4 Pole
(short)
Suitable
for coil (Cu)
Suitable for Ti
Layer 1/2 pole
Similar to
WK-09-125BG-350
There are typical gauge types being regularly reordered or anyway available at FNAL
They are suitable and enough to cover the needs for the 15 model
VMTF limitation:
At VMTF we are capable
of reading no more than 64
strain gauge channels
through the standard DAQ
(typically reading at <1 Hz)
During loading operations
there are no specific
limitations

## Slide 4: Strain gauges connections
P+ (red)
S- (wht)
S+ (grn)
P- (blk)
ACTIVE GAGE
Bonded in Axial Direction
TEMP-COMPENSATING GAGES
“Floating”
ACTIVE GAGE
Bonded in Transverse Direction
from LBNL 23N928 SPEC DWG
LBNL COIL INSTRUMENTATION
We intend to use the
typical LBNL SG
configuration where
possible
When I later refer to a
gauge (SG) I mean gauge
channel except if stated
otherwise

## Slide 5: 15 T magnet: Strain gauges on coil (L1)
Separation between
two pole blocks
Mid-plane coil SG
Pole side
coil SG
Example from 11 T models
At several inches from the pole block separation
Coil SG positions in 15 T
There are 8+4 SG per coil (24 in total) – 16 azimuthal and 8 axial.
Not all of them will be read – the ones in black ( ) are to be skipped (or back up)
The axial SG are optional and debatable ( ) – from review comments
So 8 (+ 8) active coil SG foreseen.

## Slide 6: 15 T magnet: Strain gauges on pole (L1)
Separation between
two pole blocks
Pole SG
Example from 11 T models
Pole SG positions in 15 T
There are 4 SG per coil (8 in total) – half azimuthal, half axial.
8 L1 pole SG foreseen.
SG are attached from below

## Slide 7: 15 T magnet: Strain gauges on pole (L3/L4)
Each groove will have one azimuthal and one axial gauge.
12 L3 and L4 pole SG foreseen for the magnet.
L3
L4
Groove positions circled in red show will contain each one azimuthal and one axial SG
4 SG
2 SG
Optional but
not practical

## Slide 8: 15 T magnet: Skin and bullet strain gauges
16 skin SG foreseen.
Example from 11 T models
Two positions (~ 70 cm apart longitudinally) around the magnet center to be instrumented
- At 90o and 45o
- Azimuthal and axial gages
- Two opposite sides of the magnet circumference
bullet
15 T model
Instrumented
skin gages
There are 8 bullet holes on the LE and 10 bullet holes on the RE
- A bullet is instrumented with 2 SG
- One TC on each side
- Only part to be instrumented
20 bullet SG foreseen.

## Slide 9: 15 T magnet: Strain gauge wiring
Layer 4 SG are directly wired to the trace strip pads in the grooves
Layer 3 SG are wired to the trace on Layer 4 through pole holes
Only 24 strips are available on the trace for SG
thus only 6 SG can be read during testing
Layer 2 has no SG
Layer 1 SG are wired manually (as in 11 T models)
The rest (skin, bullets) – as usual

## Slide 10: 15 T magnet: Total number of SG
 | coil | pole | skin | bullet |  | SUM
All | 16+8 | 20 | 16 | 20 |  | 72+8
Active | 8 | 20 | 16 | 20 |  | 64
We want to read
all of them
This is already a
reduced number
We can probably skip
2 SG (half the axial ones)
With this we instrument
9 out of 18 bullets; it could
be reduced to read more coil gauges
Instrument more coil gauges?
During
magnet
testing
To be instrumented
only if needed

## Slide 11: 15 T magnet: L4 grooves