## Slide 1: Wire Strain Characterization at ASC‒NHMFL
TFD Wire Specification Review
October 28, 2021
Najib Cheggour, Emsley L. Marks, William Starch, and GianLuca Sabbi
ASC-NHMFL, Florida State University, Tallahassee, FL 32310
Lawrence Berkeley National Laboratory, Berkeley, CA 94720
Work for LBNL Subcontract NO. 7560179

## Slide 2: Introduction
Work exploratory in nature, guided by the results obtained on HL-LHC RRP Nb3Sn wires.
    https://www.nature.com/articles/s41598-019-41817-7)
Ic measurements were constrained by
Power supply (limited to ~ 600 A)
Magnet (quenched at 16.5 T)
Strain Irreversibility Cliff (SIC)
Ic() “Elbow” at high strains (beyond irr)

## Slide 3
Two RRP Nb3Sn Billets
CO11S20235A04A (CERN) and 20506-4A
Four Heat-Treatment Temperatures (50 hours)
620, 665, 681, and 695C

## Slide 4
    This graph shows an example of Ic dependence on applied strain ε at 16 T to find its irreversible strain limit εirr
620  C – Determination of εirr (Billet CO11S20235A04A)

## Slide 5
    This graph shows Ic dependence on intrinsic strain ε0 (= ε – εmax ) at 16 T for two samples
    Even without normalizing Ic values to Ic-max, the three taps show very similar results throughout the strain range investigated.
    Values of εmax and εirr,0 marked on this graph are average values of those obtained for each tap
# of segments | HT | irr,0 (%) Average 0.1 V/cm | irr,0 (%) Spread 0.1 V/cm | Sigma (%) 0.1 V/cm | Ic-max (A) Average 0.1 V/cm | Sigma (A) 0.1 V/cm
6 | 620 ͦ C/ 50 hrs | 0.10 | 0.09 –0.12 | 0.009 | 467.5 | 3.1
620  C – Ic(ε) at 16 T (Billet CO11S20235A04A)

## Slide 6
    This graph shows an example of Ic dependence on applied strain ε at 16 T
    As expected, current limitation (< 620 A) did not allow measurements at all strain values.
    We could not measure εmax but estimated it and found the value of εirr
665  C – Ic(ε) at 16 T (Billet CO11S20235A04A)

## Slide 7
665  C – Ic(ε) at 16 T (Billet CO11S20235A04A)

## Slide 8
    This graph shows Ic dependence on intrinsic strain ε at 16 T for two samples.
    As expected, current limitation (< 620 A) did not allow measurements at all strain values.
    We could not measure εmax but estimated it and found the value of εirr
    Summary for the two samples, three taps each, at 16 T:
# of segments | HT | irr,0 (%) Average 0.1 V/cm | irr,0 (%) Spread 0.1 V/cm | Sigma (%) 0.1 V/cm | Ic-max (A) Average 0.1 V/cm | Sigma (A) 0.1 V/cm
6 | 665 ͦ C/ 50 hrs | 0.30 | 0.28 –0.31 | 0.013 | Maybe ≥ 680 A? | ?
665  C – Ic(ε) at 16 T (Billet CO11S20235A04A)

## Slide 9
681  C – Ic(ε) at 16.4 T (Billet 20506-4A)
    This graph shows an example of Ic dependence on applied strain ε at 16.4 T
    At 16.4 T, the window of missing data was reduced substantially.

## Slide 10
    This graph shows Ic dependence on intrinsic strain ε at 16.4 T for two samples.
    Summary for three samples, three taps each, one at 16.3 T and two at 16.4 T:
# of segments | HT | irr,0 (%) Average 0.1 V/cm | irr,0 (%) Spread 0.1 V/cm | Sigma (%) 0.1 V/cm | Ic-max (A) Average 0.1 V/cm | Sigma (A) 0.1 V/cm
9 | 681 ͦ C/ 50 hrs | 0.38 | 0.35 –0.39 | 0.013 | Maybe ≥ 640 A? | ?
681  C – Ic(ε) at 16.4 T (Billet 20506-4A)

## Slide 11
    This graph shows Ic dependence on applied strain ε at 16.4 T for two samples.
    Summary for two samples, three taps each, at 16.4 T:
# of segments | HT | irr,0 (%) Average 0.1 V/cm | irr,0 (%) Spread 0.1 V/cm | Sigma (%) 0.1 V/cm | Ic-max (A) Average 0.1 V/cm | Sigma (A) 0.1 V/cm
6 | 695 ͦ C/ 50 hrs | 0.39 | 0.36 –0.41 | 0.006 | Maybe ≥ 640 A? | ?
695  C – Ic(ε) at 16.4 T (Billet 20506-4A)

## Slide 12
The Strain Irreversibility Cliff in this RRP 162/169 wire
    This graph shows SIC for two 108/127 standard-Sn and Ti-doped RRP billets (published in https://www.nature.com/articles/s41598-019-41817-7)

## Slide 13
    This graph shows SIC for two 108/127 standard-Sn and Ti-doped RRP billets (published in https://www.nature.com/articles/s41598-019-41817-7)
The Strain Irreversibility Cliff in this RRP 162/169 wire

## Slide 14
    This graph shows SIC for two 108/127 standard-Sn and Ti-doped RRP billets (published in https://www.nature.com/articles/s41598-019-41817-7)
    Results obtained so far on two RRP 162/169 billets are also shown for comparisons. SIC might be shifted slightly to lower temperatures but cannot be confirmed for sure.
    The general trend seems similar to 108/127 RRP billets
The Strain Irreversibility Cliff in this RRP 162/169 wire

## Slide 15
-CuSn-phase Disease in samples reacted at 620 °C
Presence of the -CuSn-phase confirmed in samples reacted at 620 C
Micrograph made by Elzbieta Krekora

## Slide 16
The “elbow” is present – should HT be increased > 665  C?
    Notice the “elbow” at high values of ε0, consistent with published data on a Hi-Lumi pre-production billet (https://www.nature.com/articles/s41598-019-41817-7; Figure 8)
    Increasing HT temperature beyond 665 °C might further improve strain properties.

## Slide 17
The “elbow” disappears for HTs at 681 and 695  C

## Slide 18: Summary
Despite experimental challenges for characterizing 1.1-mm diameter RRP Nb3Sn wires, we were able to obtain strain data on samples reacted at 620, 665, 681, and 695 °C for 50 hours
Value of Ic-max increased by at least 45 % for samples reacted at 665 °C as compared to 620 °C
Value of εirr,0 increased from 0.1 % to 0.3 % by increasing HT temperature from 620 °C to 665 °C, signaling the existence of SIC. Presence of the -CuSn-phase causing the SIC was confirmed in samples reacted at 620 °C.
Further increase of εirr,0 to 0.39 % for samples HT’ed at 681 and 695 °C. Not clear if this increase is a variation from billet to billet or due to the HT temperature itself.
Ic() degradation in the irreversible regime is significantly less in samples reacted at 681 and 695 °C compared to samples reacted at 665 °C. There may be some benefits in doing heat-treatments at least 680 °C (instead of 665 °C).
General strain behavior of 162/169 RRP wire seems to follow that of Ti-doped, standard-Sn, 108/127 RRP wires