## Slide 1: Advanced Diamond Diagnostics and Detectors for Accelerators
John Smedley
Jen Bohon
LA-UR-20-20731
UC Davis
UC Fees Diagnostics Program Kickoff
1/25/2020

## Slide 2: LANL Interest
Diagnostics Challenges:
  GHz detection for photons & protons
    Flux, position, timing, IMAGING
  Expanding dynamic range - measuring the full X-FEL beam flux (in addition to pre-lasing beam)
  Shot to shot beam imaging for X-FEL beams
LANL Development Direction:
  Diamond sensors (single crystal and polycrystalline electronic grade diamond)
Partners needed for electronics development

## Slide 3: Diamond as an X-ray Sensor
Diamond Advantages:
Low X-ray Absorption
High Thermal Conductivity
Mechanical Strength
Radiation Hardness
Indirect bandgap
Sample Information:
Electrical grade CVD single crystal diamond
  (100) surface orientation
  ~1 ppb nitrogen impurity
Typical size: 4mm x 4mm x 50µm
Low energy x-ray
High energy x-ray
X-ray generated charge carriers
Electron hole pairs created near incident electrode: must move entire thickness of the diamond
Electron hole pairs created throughout the thickness creating a column of electron-hole pairs

## Slide 4: Diamond Sensors
Diamond Applications:
Synchrotron beam diagnostics
  White and Monochromatic
  Nondestructive, High power
  Flux, Position and Timing
  Beam Imaging
Medical beam diagnostics
  Gamma Ray Therapy
  Proton and Ion beam Therapy
FEL beam diagnostics
  Promising – current R&D effort
Diamond Properties:
Low X-ray Absorption
High Thermal Conductivity
Mechanical Strength
Radiation Hardness
Indirect bandgap (no radiative recombination)
Full charge collection at modest bias (0.1 V/µm)
Fast - carrier mobility 4500/3800 cm2V-1s-1 (e/h)
Flux linearity over more than 11 orders of magnitude!
R&D100 2016 Finalist
Zhou et al. (2015)
JSR 22 :1396
(cover article)
1k pixel transparent camera
Sensors in use at synchrotron beamlines around the world

## Slide 5
“Clean” Single Crystal
300 µm thick plate
Polycrystalline
210 µm thick plate
Response vs Flux and Bias
Response to incident flux linear over 11 orders of magnitude
Under 0.1 V/µm required for full collection

## Slide 6
Wire Bonding - Instrumentation
25 µm aluminum wirebonds
5 bonds per pad
Conductive epoxy for backside/bias contacts
3.1mm
4.0mm
20µm
75µm
Lithography @ CFN
Electronic grade single crystal (100) diamond
30-50 µm thick
20 µm street over a 1mm center region
Metalization: 25 nm Pt
Topography – NSLS/CHESS
White beam topography
Prior to slicing
Fabrication

## Slide 7: Beam Position Monitors
4mm
3.5mm
50 µm
A
C
B
D
RMS ~ 38 nm

## Slide 8: Thin Detectors
RIE ultrathin layers
4 µm membranes
Transmits 2 keV photons
Potentially much faster response
4 devices in use at SMI
5V

## Slide 9: All Diamond Detectors
Patterned UNCD Contacts
Diamond Contacts – no longer the weakest link
UNCD Grown at ANL
Lithographically masked RIE
No fluorescent lines

## Slide 10
Conduction comes from sp-2 content (along grain boundaries)
Flux linearity limited at very high flux (>1013) due to lower contact conductivity
J. Synchrotron Rad. (2018). 25, 1060–1067

## Slide 11: Transmission-mode Diamond Imaging Detector
Pixels are created by metalizing one side of the diamond with horizontal stripes and the other with vertical stripes. As the x-rays pass through the diamond, the induced current is collected in each vertical stripe, while the bias is applied to individual stripes on the other side. This bias is cycled, allowing readout of one line of “pixels” at a time.
Image Readout
32 x 32 stripes, yielding 1024 pixels
Only one row is active at a time minimizing ohmic heat generation.
Imaging at 32 Hz
Dynamic range from 50nA to ~10mA per “pixel”
Applications
Freestanding detector provides beam shape, flux and position during optics adjustment and/or during experiment
Imaging window: diamond sensor brazed to a stainless steel vacuum flange serves as both window and beam diagnostic
readout animation

## Slide 12: Beam Imaging
CHESS G3 Beam
NSLS X28C Beam

## Slide 13: Beam Imaging
R&D100 2016 Finalist
Zhou et al. (2015) JSR 22 :1396 (cover article)

## Slide 14: Imaging Detector at NSLS-II XFP
Beam Images at Different Mirror Settings
Center for Synchrotron Biosciences
Integrated into XFP Beamline Controls System (EPICS/CSS)
Used to align XFP toroidal focusing mirror (beamline commissioning)
Pitch: 4.78 Focus: 2.6
Pitch 4.72 Focus: 3

## Slide 15: Polycrystalline Diamond Imaging Detector Results are Promising
ZW2: 50 µm thick Polycrystalline “Electronic Grade” Diamond (low nitrogen content) - material and contacts fabricated by Applied Diamond, attached to PC board with high-vacuum epoxy
Cr/Au contacts, with ~400 nm Au (very thick!)
200 µm pitch for “pixels”, only 28/40 wire bonded (5.6 x 5.6 mm2 active area)
Bias scan – 1 mm pinhole (leakage on the order of 100 pA) – photoconductive under positive bias, but nicely saturates at ~ full collection (similar to previous E6 poly) in negative
Flux scan – 1 mm pinhole (results still in analysis)
Response map (uniformity) – 30 µm pinhole
Image – used 1 mm diameter pinhole to define the beam (of ~ 3 mm beam)

## Slide 16: Application to FELs
Can not rely on pulse to pulse stability
Can get shot by shot line out by measuring current on all detector channels
Need to redesign electronics to allow FPGA to measure current and/or bias strips independently
Can also be configured as a variable-center quad, expanding the spatial dynamic range of the device for small beams

## Slide 17
Ring mode “hybrid fill, top up”.
102mA total, 16mA in first bunch, 86mA in remaining pulse train.
Separated by 1.594 µs
Ratio of ring currents matches very closely to measured charge ratio
Current Ratio: 86mA/16mA = 5.38
Measured Q Ratio: 0.91nC/0.17nC = 5.35
Pulse Mode Beam Position 11-ID-D, APS

## Slide 18: NSLS-II Properties and Capabilities for Detector Calibration and Testing
NSLS-II Ring Parameters | 
Energy | 3 GeV
Current | 500 mA (400 mA currently)
Pulse length | 30 ps
Time between bunches | 2 ns (500 MHz)
Average bunch current/charge | 0.47 mA/ 1.25 nC
# of buckets filled/total | 1056/1320
Note: one timing pulse available in ion cleaning gap, NSLS-II working on varying up to 5 mA in that pulse; period of 2.64 µs. | 
XBIC (X-ray Beam Induced Current) Endstations
2 endstations (ISS 8-ID, XFP 17-BM) for detector calibration at NSLS-II
ISS: Damping wiggler source; Mono (4.9 – 36 keV), up to 1014 ph/s in ~0.5 x 0.5 mm beam
XFP: Focused white beam (5 ~ 16 keV), up to 2x1016 ph/s in variable 120 µm x 450 µm to 2.7 mm x 2.7 mm beam
Capabilities
3 x 3 x 3 ft3 nitrogen environment with cable pass-thru (or old style 1 x 2 x 2 ft3 compact “fishtank” if space is limited)
Integrated motion control for XBIC mapping, optical mounting
Multiple calibration options (Ion chamber, diamond detectors, copper block calorimeter)
Pinholes 25 µm and up, attenuators
Pulsed or standard DC biasing
Current measurement from pA to 100s of mA
Unique prototype electronics and epics/python-based readout and controls for diamond imaging detector
Can attach fast scope for timing measurements
Funded through a BNL LDRD project

## Slide 19: Proton Irradiation at LANSCE
LANSCE blue room setup
Proton Energy: 800 MeV
Beam Flux: 1011 proton/(cm2·s)
Beam Size: diameter ~ 2.5 cm
Exposure Time: 2180 seconds
Incident Energy: 4.486 kJ
Electrical Field: +0.1 V/µm
Incident Proton Beam
#1
#3
Real-time IBIC Monitoring
#2
Charge collection in the material changes – 20% loss in efficiency with 10yrs of dose
Damaged detectors retain linearity with flux (but with longer charge collection distances)
Response Linearity
W=13.4±0.04 eV
I-V curve
0.6 V/µm
Performed at WNR target site: additional LANSCE capabilities available

## Slide 20: Diamond Detector Research Program
Proton/ion beams
  Improved diagnostics for LANSCE
  Improved targeting of beam on patients for medical dosimetry (DOE HEP)
  Anticipate same spatial and temporal resolution as for x-rays
FELs
  Improved beam targeting (< beam size)
  Achieve linearity over full gain curve (104 – 1011 ph/pulse)
  Diagnostics for future FELs (DMMSC) enabling full gain curve optimization
  Beam profiling: shot to shot profiles (flux, position, horizontal and vertical distribution)
External case (similar)
Circuit board design
(Integrated current)
Response linear to ~3% of full beam
Energy: 6080 eV
LANSCE (Target 4) Proton Beam Measurements
Timing measurement reproduces LANSCE pulse structure
Lifetime measurement (100 years of therapeutic dose) – 20% reduction in efficiency after 10 years of dose
Looking for collaborations to continue this development and calibration effort
LANSCE ICT
Flux Linearity (SwissFEL Bernina Measurements)
(Shot-by-shot)
Individual pulse measurement, pulse broadening beyond linear regime
Beam attenuation

## Slide 21: LANL Work Packages
WP 1 – diamond device fabrication
  Sensor for full FEL gain curve
  Sensor for FEL beam imaging
  Sensor for proton imaging
  Sensor for Dynamic materials science
WP 2 – Device testing (not just diamond)
  NSLS-II/CHESS
  SwissFEL/LCLS
  LANSCE
LoE – Jen (0.25), Postdoc (0.5), 20k travel – 200k/yr
UC support – materials purchase (diamonds) – 20k/yr, student participation?

## Slide 22: Expanding the scope of photophysics beyond photoemission: x-ray detector development
Goal: to achieve linearity over full gain curve (104 – 1011 ph/pulse)
  Improved beam targeting (< beam size)
  Diagnostics for future FELs (DMMSC) enabling full gain curve optimization
  Beam profiling: shot to shot profiles (flux, position, horizontal and vertical distribution)
So far: Modified detector with additional capacitance, achieves linearity up to 3% of full beam
  Detector performance unchanged by FEL beam (no damage)
  Limit caused by geometry and charge pileup – we have engineering solutions in mind
Response linear to ~3% of full beam
Flux Linearity (SwissFEL Bernina Measurements)
Looking for collaborations to continue this development and calibration effort
Individual pulse measurement, pulse broadening beyond linear regime
External case (similar)
Circuit board design
Diamond Properties
Low X-ray Absorption
High Thermal Conductivity
Mechanical Strength
Radiation Hardness
Indirect bandgap (no radiative recombination)
Fast - carrier mobility 4500/3800 cm2V-1s-1 (e/h)
J. Bohon, P. Juranic, E. Muller, M. Zou, A. Sumant, U. Jastrow, K. Tiedtke, T. Tanaka, S. Owada, R. Follath, C. Arrell, Y. Arbelo, M. Degenhardt, A. Sorokin, J. Smedley

## Slide 23: Diamond Beam Imager
Development of a diamond window which will provide position, flux and morphology of high flux x-ray beams while simultaneously acting as the vacuum-air interface.
Major Challenges for high flux beamlines
Heat load management in optics (including beryllium windows)
Real-time volumetric measurement of beam properties such as flux, position, and morphology.
X-ray footprinting (XFP) at NSLS-II
Focused white beam
Variety of beam sizes/shapes needed
Feedback and control systems for optical elements or sample positioning stages.
Image Readout
32 x 32 stripes, yielding 1024 pixels
Only one row is active at a time minimizing ohmic heat generation.
Real-time imaging at 32 Hz, FPGA control
Allows for feedback control of beamline components.

## Slide 24
Chart: Time (ns)
Temporal Response, Hard X-rays
J. Bohon, E. Muller and J. Smedley, J. Synchrotron Rad. 17, 711-718 (2010)

## Slide 25
Chart: Time (ns)
Temporal Response, Soft X-rays

## Slide 26: Radiation Damage effect on Temporal Response
NSLS II synchrotron beam timing
After 1014 protons (100 years of therapeutic dose), detector bulk response ~30% of initial
Carrier property (at 400V):
Transit time: ~ 0.67 ns
 carrier velocity ~1209 cm2/(V∙s)
Trapping time: ~ 4.18 ns
=> Charge collection is significantly delayed by radiation damage
Real-time IBIC Monitoring @0.1 V/µm
Sample #1
Sample #2
Sample #3
Fitting Curve #1
Fitting Curve #2
Fitting Curve #3