## Slide 1: Potential SLAC Collaboration
25 January, 2020
Bryce Jacobson – Diagnostics & HLA Group Leader (AD-LFD)
Phil Heimann – Matter in Extreme Conditions (LCLS)
Alan Fisher – Diagnostics & HLA Group (AD-LFD)
LCLS, SSRL, UED, Linacs…

## Slide 2: Sharing the SLAC Linac
  Linac Tunnel and Klystron Gallery, 3 km 1962–
  LCLS, the first hard x-ray FEL, 3rd km 2009–2018, 2019–
  FACET, advanced acceleration tests, 1st and 2nd km 2010–2016
  LCLS-II, 4-GeV superconducting linac, 1st 700 m 2021–
  FACET-II, 2nd km 2020–
  LCLS-II HE, 8-GeV upgrade, 1st km ≈ 2026–
Soft x-ray undulator
FACET
LCLS
1-km linac
Beam transport
Soft x-ray undulator
LCLS-2
Bypass beamline
Hard x-ray undulator
Users
FACET-2
SLAC linac
3 km
Superconducting linac
RF gun
RF gun
Thermionic
gun
RF gun
1-km linac
2-km linac
Experimental area
Users
LCLS-2 HE
Beam switchyard
Bypass beamlines
Undulator
Users
Experiments
Beam transport

## Slide 3: X-ray FEL Source Fluctuations
Spatial
Spectral
Temporal
Every pulse is different and must be diagnosed individually
Can be done “off-line” in analysis (XTCAV, TimeTool)
Two linacs – 120 Hz copper and CW SC
Up to ~ 1 MHz rates – or any structure at 1 MHz
Energy Recover Linacs up to 1.3 GHz
Future XFEL oscillators
Stabilize the accelerator!! – Feedbacks, Optimizers, Machine Learning, “Virtual Diagnostics”…

## Slide 4: LCLS-II Diagnostics Challenges vs. LCLS Solutions
Repetition rate 0.9 x 106 Hz. Power up to 200 W.
Instrument | TMO | TXI | 2.2
Photon energy (eV) | 250 - 2000 | 250 – 7000 | 250 – 1600
Pulse energy min (μJ) | 10 | 30 | 0.2
Pulse energy max (mJ) | 12 | 8 | 0.03
Repetition rate max (Hz) | 0.93x106 | 105 | 0.93x106
Device | Primary Function
Profile Monitor (PM) | X-ray beam alignment
Gas Monitor Detector (GMD) | I0
Fluorescence Intensity Monitor (FIM) | I0
Arrival Time Monitor (ATM) | X-ray / optical laser timing
Wavefront Sensor | X-ray optics alignment Focal spot characterization
X-ray Spectrometer | X-ray spectrum
Gas Monitor Detector
Arrival Time Monitor

## Slide 5: Profile Monitor & Power Meter
Purpose: X-ray beam alignment, to observe the centroid and profile of the X-ray beam.
Power requirement: operate the accelerator at a low repetition rate or use the gas attenuator in the front end enclosure.
The average X-ray intensity along the beamline will be measured by power meters from Gentec, positioned below the YAG crystals.
Ce:YAG is not compatible with power > ~1 W.
Power Meter

## Slide 6: LCLS Detector: ePix 10k
 | ePix 10k Quad
2D pixel array detector | 0.5-1 mm thick Si
Pixel size | 100x100 µm2
Maximum range | 10,000
Readout | 5-120 Hz, single shot
Tile size/ total active chip area size | 40x40 mm / 80x80 mm, Gap ~similar to cspad
X-ray coverage | 4-24 keV

## Slide 7: Point and Long Beam-Loss Monitors
  PBLM Point Beam-Loss Monitor
    Placed at likely loss sites
    To prevent damage to collimators, stoppers, dumps
  LBLM Long Beam-Loss Monitor
    To span the full machine in segments, from gun to beam dumps
      To detect losses in less likely places
      To identify loss locations
    To protect people
      In the Klystron Gallery, above penetrations from the linac tunnel
      Outside the linac-to-undulator (LTU) transport line
        A concrete bunker that is above ground and so less shielded than the linac tunnel

## Slide 8: LCLS-II Beam-Loss Monitors Serve Three Functions
  BCS Beam Containment System
    Stops the machine if beam goes outside its intended location or power
    Simple trip logic, without software, without knowledge of timing or rate
  MPS Machine Protection System
    Halts or rate-limits a beam path if losses exceed the threshold for machine damage
    Trips before BCS: Threshold generally set 10 times lower
    More intelligent: Uses software, knows beam timing and beam paths
  Diagnostic Signals
    Help operators to tune: Display losses below trip threshold, indicate loss location
    Reduce radiation on beamline components and outside the tunnel
    Detect loss signals from fast wire scanners
  EPS Equipment Protection System
    Not implemented (see FE monitors)

## Slide 9: Beam-Loss Requirements for LCLS-II
  Loss detection at SLAC had always used ionization chambers, but:
    Ionization chambers can have a nonlinear response to high loss power (IBIC 2016)
    High losses at BCS thresholds: 100 W for PBLMs, and 35 to 1000 W for LBLMs
    Dynamic range: Continuous RF may cause field-emission with significant power
 | LCLS | LCLS-II | LCLS-II HE | 
Repetition rate | 0.12 | 929 | 929 | kHz
Allowed beam power | 0.5 | 120 | 250 | kW
“Maximum credible” power | 0.5 | 1100 | 2200 | kW
Loss integration time | 1 | 0.5 | 0.5 | s
Trip time | 8.3 | 0.2 | 0.2 | ms

## Slide 10: A Diamond Detector as a Point BLM
  Metallized diamond chip, 4×4×0.5 mm3 for the single-crystal (B1) type
  Radiation creates electron-hole pairs, which are collected by a bias voltage
    Electrons and holes have high mobilities and cross the thin chip quickly
      Like a solid-state ionization chamber, but fast: Few-ns pulses
  Large dynamic range
    1 to 106 MIPs (minimum ionizing particles)
  Unlike ionization chambers (SLAC’s old PBLMs):
    No pile-up from prior pulses
    No gas: low maintenance, high reliability
Cividec B1HV

## Slide 11: Polycrystalline versus Single-Crystal Diamond Detectors
  ≈1 W of beam (<< MPS trip level) scraped at energy collimator (LCLS BC2)
  Beam rate quickly toggled between 0 and 120 Hz
  Cividec B2, polycrystalline
    18-s decay tail
  Cividec B1, single crystal
    Much better: No tail
  Also, first test of heartbeat
    Modulating B1 bias voltage

## Slide 12: Modeling the Diamonds: Area Coverage
  Protection collimators are used after the linac and transport lines, where two or more beamlines are close together
  Extension to FLUKA computes B1 diamond signal from a loss shower
    Uses Cividec data on creation of e-h pairs
  Signal is nearly constant downstream of loss on protection collimator, for >20 m
    Diamond is 1.5 to 2 m from beamline
    Experimental test in End Station A
  Each diamond can cover an area, rather than a single protection collimator
Beam-Loss Model
Protection collimator
Beampipe

## Slide 13: 2017-01: Field-Emission Test on CEBAF at JLab
  Last cryomodule in the north linac of CEBAF (Jefferson Lab)
    1L26, a high-gradient C100 type
  RF was on in all linac cryomodules, but without photocurrent
    Cavity gradients in 1L26 were raised sequentially in steps of 1 MV/m
    RF in other modules held constant
  Fiber with PMT:
    12 m of fiber strung along the side of the tank
    Uncooled blue PMT placed at upstream end and shielded with lead
      Before the red-sensitive, cooled PMT was selected
  Diamonds:
    Diamond detector at each end of the tank, above warm beampipe
      A known loss point, since the warm pipe has a smaller diameter
  BLMs:
    JLab BLMs are PMTs (without scintillators), at each end of the tank.

## Slide 15: 2017-01: Stepping CEBAF Cryomodule Gradients
  Fiber displays growth of field emission with increasing cryomodule gradients
  Fiber is far more sensitive to the inner cavities: spike in cavity 7, drop-out of cavity 3
  Downstream diamond (high-rad B4 types, less sensitive) and BLM show steps of cavity RF
  Upstream BLM and diamond dominated by losses from upstream (1L25); weak response to 1L26
Cavity Gradients (MV/m)
Loss Signals
Steps of 1 MV/m
Upstream BLM
Downstream Diamond
Upstream Diamond
Downstream BLM
Fiber
Cav 7
Cav 3
Cav 8